## Slide 1: Improving Precision and Accuracy for Trace Elements in EPMA
J. J. Donovan University of Oregon, CAMCOR, Eugene, OR 97403-1253

## Slide 2: Traditional Off-Peak Methods
Bgd “low”
Bgd “high”

## Slide 3: Two solutions: MAN and MPB
MAN (mean atomic number)
  Measure the “on-peak” background in (standard) materials that do not contain the element of interest
  Correct for continuum absorption and fit to average Z
MPB (multi point background)
  Measure the off-peak background at more than two positions
  Fit to polynomial or exponential expression

## Slide 4
DyLb1 peak
FeKb
MPB (multi-point backgrounds) provide accurate trace analysis background measurement, especially in complex systems
Jercinovic et al. 2012, Mater. Sci. Eng. 32, 012012

## Slide 5
List of datasets
Measured elements, Ag La selected
wavescan overlay
peak intensity
background intensities (circled – currently used for fitting)
background interference
data rows in dataset
x-ray line markers
Manual or automatic optimization, with many useful features for trace element evaluation

## Slide 6: ”Shared” MPB
?
?
?
?

## Slide 7
By prior determination of continuum intensities in a variety of materials (over a range of average atomic number) that do not contain the element of interest, the matrix iteration can quantitatively calculate the background intensity from the unknown specimen composition.
Since no off-peak measurements are utilized, time is saved resulting in less sample damage and elimination of possible off-peak interferences.
This is a significant acquisition time advantage for major and minor elements.
… and also results in improved trace element sensitivity…
Donovan, et al., Amer. Min., 2016
MAN (mean atomic number) background calibration

## Slide 8
Comparing MAN and Off-peak Measurements (same dataset)
No Blank Correction
Blank Correction

## Slide 9: On and off peak variances add in quadrature!
When we directly measure the background intensities, we introduce additional imprecision to the net intensity calculation.
If only there was a method to determine the continuum background intensities *without* a direct measurement…
How about the MAN background method?

## Slide 10
Modeling Precision (Sensitivity) with the MAN Background Method
MAN background intensity is correlated with Z-bar (the average atomic number) in the form of a line with slope (m) and intercept (b)
The MAN background variance is therefore
For the linear case, the partial derivatives can be combined to yield:

## Slide 11
Calculating the variance of the average atomic number, Z-bar:
or
The total MAN variance is therefore the average Z variance (for measured elements) plus the regression variance, correct?
“Model A”

## Slide 12
 | Th | Hf | U | P | Y
Average | .226061 | 2.02225 | .611455 | .320043 | .113804
Std Deviation | .028376 | .142604 | .128880 | .037488 | .021821
Off-peak Model | .027325 | .081837 | .044766 | .032520 | .019335
 | Th | Hf | U | P | Y
Average | .254557 | 2.04241 | .630875 | .280978 | .113875
Std Deviation | .000348 | .001509 | .002499 | .000904 | .000156
MAN Model (A) | .025887 | .076575 | .026831 | .017200 | .006861
MAN Model (B) | .000328 | .001510 | .000612 | .000265 | .000095
Off-Peak Background Correction (calculated background pixels)
MAN Background Correction (fixed ZrSiO4 matrix)
 | Th | Hf | U | P | Y
Average | .254586 | 2.04253 | .630940 | .281020 | .113894
Std Deviation | .000812 | .002682 | .002897 | .001754 | .000683
MAN Model (A) | .027446 | .083754 | .029741 | .018462 | .007315
MAN Model (B) | .001886 | .008685 | .003521 | .001526 | .000549
MAN Background Correction (measured ZrSiO4 matrix)
Cps/nA

## Slide 13
“Model B”
MAN Variance Models
“Model A”
MAN regression precision is actually an accuracy issue!*
Unless we remeasure our MAN curve for *every* point or pixel!

## Slide 14
Natural Quartz
Quant mapping

## Slide 15
Natural (Toba) Zircon
~20 hours
~10 hours

## Slide 16
Off-peak measurements =
worse sensitivity in twice the time!

## Slide 17
14.8 wt% (publ.)
20 nA, 0.5 sec per pixel